.Small, fine TransistoCurveTracer

Why do you need a TransistorCurveTracer („TCT)“ ?

Isn't a transistor tester or a voltmeter enough?
A transistor tester is very helpful for a first check, especially because it can also determine and display the type of transistor and the pin assignment.
t.
But the output values only refer to one operating point (collector current).
For the circuit developer, however, it is often important how the component reacts to different
working points - e.g. base currents - behaves.

What is the difference between the TCT presented here and others?

     • Compact construction: dimensions: 8x9x6cm , weight: < 500 g (incl. 18650 Lio battery).
     • independent of external power supply (battery operation, with a 18650 cell)
     • Designed for bipolar transistors/diodes, MosFets, J-Fets(incl. Depl.Mode MosFets).
     • Collector/Emitter voltage < 12V, 0 ..150mA optimal for small power transistors
     • Base current of approx. 5µA ... 1000µA with a precise constant current source,
     • Gate voltage 100 mV ...12 V in 100 mV steps (changeable),
     • Adjustable current limitation (e.g. diodes 5mA, 50mA,...)
     • relatively large color display (3.2")
     • easy to use and fast input via rotary encoders on interrupt pins
     • high accuracy and clear characteristics thanks to 12 bit ADC's / DAC's
     • Various scales for x (U) and y-(I) axis: detailed display of interesting areas
     • Direct comparison of characteristics (types of the same polarity or complementary type)
     • Display of typical values such as Beta(Ib), V-Threshold,V-Pinch, Idss
     • Bluetooth interface (BLE 4.0) for data transfer to the PC or smartphone/tablet
     • Data transfer via USB to display the curves on a PC with e.g. „Processing“.
     • The input parameters are saved in the internal EEPROM
     • Low power consumption: Uce (15V) only switched on during measurement.
     • Additional function: adjustable constant voltage source 1 mV to 12.00V, > 10 mA
     • Additional function: adjustable constant current source from 1µA to 4000 µA
     • Additional function: Step generator for low frequencies +4V, ...0V (...-4V)
     • housed on 5 boards (Kicad) with one layer
     • relatively low component costs

This TCT is based on the very interesting and well-documented project by Peter Balch: Transistor Curve Tracer [1], from which the graphics library, among other things, was taken over in a slightly modified form.

But then I moved quite far away from it - also in terms of concept - in order to get a better and lower-noise display with a significantly extended measuring range, simple operation and extended functionality, which unfortunately is also reflected in the price.

I have eliminated the previous operation via the touch panel and do the program control and parameter changes via two rotary encoders that are connected to the two external NANO's interrupt pins, allowing quick response. The input is much more pleasant and faster than by touch, especially with the small display and the many parameter values.

The replica is more complex because of some SMDs and the wiring of some boards and therefore not necessarily suitable for "beginners". Considerable procurement times (in Germany) are/was to be expected for some special ICs.

The basis is an Arduino-NANO(clone), preferably with a pre-installed Opti-Boot Loader [2,3,4], because there you have 1.5 kB more free flash memory, otherwise you only some 100 bytes of free flash memory.
(ATMega328 board with pre-installed Opti-Boot-Loader: e.g. Arduino-UNO,Adafruit Metro Mini 328, or a board with more flash and port pins (for extensions): ATMega2560 ).
(Annotation:
The UNO board or Adafruit Metro-Mini can also be used here, since I don't use the additional purely analog pins of the NANO (A6,A7).

The Adafruit board Metro Mini 328 not only has the better arrangement of the pins (all digit.pins and thus also pin13 (SPI-SCK) on one side) but also the possibility of setting the level to 3.3V with a solder bridge, so that the display and the HM-10 BLE module do not require a voltage divider for level adjustment.

The microprocessor controls two 12-bit DACs (in the MCP4822 SPI) for the Uce supply, as well as a 12-bit DAC (MCP4725-I2C) for the base/gate control, which is connected via a DC/DC converter (5 V → 30V) and is operated isolated for the I2C bus via an ADUM1250.

This totally isolated DAC (range from 1µA to 4mA, 1µA steps, used here only from 5µA to 1200µA for the bipolar transistors) generates the base current or the gate voltage in the polarity that matches the component / "device".

The base current (or the gate voltage) is switched on via an electronic switch (“Photomos relay”) with a low on-resistance and leakage curren( e.g. TLP222A).

In addition to these DACs, a 12-bit ADC (MCP3402 SPI) is also used to measure the collector-emitter voltage and the voltage drop across the measuring resistor RS (50 ohms) (to determine the collector current). It is also used to monitor the battery voltage. Currently only 3 of the 4 inputs of the ADC are used.
In addition, the NANO supplies the up to 3.2-inch color display (ILI9341) (on SPI bus).
There were some problems with the 3/4 SPI devices because of the different speeds, since the display runs at 8 MHz, the DACs at 3.6 MHz and the ADC but only at 2 MHz (and the touch is even slower, but is not used in this version) . As a feasible / not quite perfect solution, I save a maximum of 320 measured values per set base current in a table with an SPI speed of 2 MHz. Then I increase the SPI speed to 8MHz for the display and output the table values on the display, in order to then save the next 320 measured values for the next base current - with reduced SPI speed. The voltage Uc is generated across the two DACs of the MCP4822, with DAC-B only outputs the values (0V, 2V,4V ...10V) at the opamp TCA0372, DAC-A generates the intermediate values of 0.0mV, 0.5mV, 1.0mV...2000mV. They are then added to the values generated from DC-B (two inverting amplifiers cascaded), so overall the range of 0V to 12V in 0.5mV steps can be generated with very good linearity.
This resolution is a bit exaggerated, but with small voltage values (e.g. characteristic curve of Shottky diodes or in the initial area of the curves of bipolar or field effect transistors) it is useful, and the mean value of the measurements for the same x_value is also formed through "oversampling".
The TCA0372 summing amplifier with gain +1 for DAC_A and gain +10 for DAC_B is sketched in Figure 1. Since the TCA0372 is not a rail-to-rail amplifier, it needs a negative auxiliary voltage: -5V is generated from the +15V using a small voltage converter (LMO78-5-.5). With this you get clean curves even near the zero point, especially the 1, 2, 3V scales benefit from it.

  Power supply

The TCT is supplied by a single Lion battery cell (18650), the voltage for the NANO board (7.5 V) is generated with a step-up converter (MT3608). A second MT3608 is used to supply the TCA0372 amplifier board with 15V. The MT3608 has the advantage that it works from about 2.5V (important, since there is only one Lion cell is used (approx. 3V .... 4.2V). In addition, it works at around 1.2 MHz, so that there is little interference at the low frequencies.
 Its setting/commissioning is somewhat critical, since the wiper of the 10-speed trimmer is connected directly to the FB (feedback pin 3) and at the upper stop the output voltage is connected directly to FB, which is probably only 0.6V? tolerates. So, before applying the input voltage, you should bring the controller a few turns from the upper stop position (namely after higher output voltage values, i.e. turn to the left Chinese logic ??) [5,6],

To save energy, the 15V supply is only activated when the curves are processed. After switching on, the 7.5V are immediately available for the microprocessor board, which includes the display and the constant current source, which, however, requires less than 1W. The NANO Board, including the display, requires approx. 400 mA during operation, but more than 1 A when switched on. The Adafruit board needs significantly less at 3.3V because of a better regulator, and about 6V at the input is also sufficient here.

 After this technical overview, a little more about the test procedure:

 After switching on, the start screen appears on the display, on which only three large icons appear on the left side with selection fields (buttons) for:
•
    bipolar transistors (including diodes) •
    Mosfets •
    J-Fets (incl. n depletion mode mosfets)

which are controlled via a rotary encoders. A button is activated by turning the 2nd encoder. Two buttons ("P" or "N") for starting the measurement process are then displayed in the upper area.
 Before doing so, however, you should examine and possibly modify the parameters displayed in two columns below. The parameters in the middle of the display:

  • min: minimum value (base current or gate voltage) with which the measurement process is to be started
  • incr: Increment with which the base current or gate voltage is to be increased
  • max: maximum base current or gate voltage curve to be displayed
The base currents are entered in µA, the gate voltage in 100mV steps.
A diode is assumed if the field incr: is set to 0 for "bipolar".

The following parameters are shown on the left edge of the display:

   • Bluet/CCS: Transfer of data with BLE 4.0 to PC or smartphone, or constant current/constant voltage operation
   • x-scale[V]: length of the x-scale ( 1V, 2V, 3V, 4V, 6V, 12V)
   • y-scale[mA] height of the y-scale ( 2 mA, 5 mA, 10 mA, 20 mA, 50 mA, 100 mA)

Annotation:

Field Bluet/CCS = 1 Output of characteristic curve data to smartphone via Bluetooth
        = 2,...6 switching to constant voltage source, constant current source or stair step generator.
The management of the parameters is similar to that described for the start screen, you jump between the fields by turning encoder 2, selecting a parameter with the other encoder 1, changing the parameter by turning the encoder, saving/exiting the parameter with encoder 2. If you are satisfied with the settings, start the measurement process / curves by turning the other encoder: "P" (for p-devices) or "N" (for n-devices) r Here you have a short waiting time to activate, in which you should check the parameters again. During this waiting phase, the activated fields P and N flash. Only now do you see the selected scales and the selected device including icon in the upper area. The curves are built up from the x-0 point and additional data (such as beta/hfe for npn/pnp and the base current/gate voltage) are entered at the end of the curves. At the end of the measurement, the screen with the displayed curve stops and a query / request appears at the top left,

  • whether you want to end the measurement ("Stop ?")
  • whether you want to continue the measurement with changed parameters ("Param ?")
  • whether you want to compare the currently displayed component with a component of the same polarity ("Compare ?")
  • whether you want to compare the currently displayed component with a component of a different polarity (“Complem. ?”) (complementary type, e.g. anti-parallel diode)

 Annotation:
 For example, if you want to compare the characteristics of an n-device with a p-device, you first determine the characteristics of the n-device normally, exchange the n for the p-device in the text tool measurement version and start the comparison by selecting the field "Complementary “ with encoder and activation with button. The characteristics of the p-device are immediately shown in a different color on the display ("complementary type"). Components of the same polarity are compared by selecting the "Compare" field.

     Storage of changed parameters in the internal EEPROM

 The specified parameters are stored in the internal EEPROM when the curves are started, so that the last parameters entered are displayed when the device is switched on. The following addresses are assigned in the EEPROM:

    • 0: transistor type 0: bipolar,
    • 1: Bluetooth : 1=on, 0=off
    • 2: x-scale: 1,2,3,4,6,12
    • 3 y scale: 2.5,10,20,50,100mA
    • 4: minimum base current bipolar transistor
    • 6: maximum base current bipolar transistor
    • 8: increment base current bipolar transistor
   • 10: minimum step size for input min: (field 4)
   • 12: minimum step width for input max: (field 6)
   • 14: minimum step width for input incr:: ( field 8 )

   • 20: Transistor Type 1: Mosfet
   • 21: Bluetooth : 1=on, 0=off
   • 22: x-scale: 1,2,3,4,6,12
   • 23 y scale: 2.5,10,20,50,100mA
   • 24: minimum gate voltage Mosfet
   • 26: maximum gate voltage Mosfet
   • 28: increment gate voltage Mosfet
   • 30: minimum step width for input min: (field 24)
   • 32: minimum step width for input max: (field 26)
   • 34: minimum step width for input incr:: ( field 28 )

   • 40: Transistor Type 2: J-Fet
   • 41: Bluetooth : 1=on, 0=off
   • 42: x-scale: 1,2,3,4,6,12
   • 43 y scale: 2.5,10,20,50,100mA
   • 44: minimum gate voltage J-Fet
   • 46: maximum gate voltage-J-Fet
   • 48: increment gate voltage J-Fet
   • 50: minimum step size for input min: (field 44)
   • 52: minimum step width for input max: (field 46)
   • 54: minimum step width for input incr:: ( field 48 )

You can also store the device name (e.g. BC548C or other information) in the internal EEProm from address 60:

  • 60: length of the text field
  • 62 text 1st byte
  • 63 text 2nd byte • ..

. the text field / device-name is shown in the display below the P / N buttons and can be managed via a small encoder-based editor - a maximum of 11 entries. :

Data transfer via Bluetooth (BLE 4.0)

An HM-10 module with BLE 4.0 is integrated into the device, it is connected to the
serial interface RX / TX and thus to USB via a rocker switch.

The transfer is activated in the selection menu
at the top right below the "N" button. Bluet./CCS: 1 (0 for Bluetooth not activated

the following values are transmitted with Separator Linefeed (“\n“ ):

  • X: x scale and then 1,2,3,4,6 or 12 the voltage end values of the x-axis
  • Y: y-scale and then 2,5,10,20,50,100 the current end values of the y-axis
  • P or N : the polarity (npn/pnp,nMos,pMos,nJFet,pJFet)
  • T, M, F (D) : the type (transistor(bipolar), mosfet, jFet (diode not real yet.)
  • C: Curve per base current or gate voltage Here, up to 320 values of the y-axis - separated by commas - are output
  • Q: end of curve • (not activated)
  • B: or G : for base current or gate voltage
  • E : Error Error message when the maximum collector current is exceeded
  • H : hfe/beta of bipolar transistors as a function of base current (at the end of the curve)
  • O: Ohm (at j_Fet with scale 1V determined at Uds =0.5V)
  • D : diode current at Ud in mA (diode is "recognized" if at npn/pnp incr=0)
  • U : diode voltage Ud at Idiode in mV (flux or blocking voltage at Zener < 12V)
  • Z : end of curve

Note When the HM-010 module (BLE 4.0) is switched off, in the position Bluet./CCS = 1
 The above parameters can be transferred to the PC via USB and monitored with TeraTerm, for example.

  Use as a precise constant voltage source

The mode: constant voltage source is activated with Bluet./CCS = 2:
 here you can use the parameters min.incr and max: the voltage value specify up to a maximum of 12 V.

 Use as a precise constant current source

 The mode: constant current source is activated with Bluet./CCS = 3: Here you can enter the constant current in µA (1...4000) in the max: parameter.

Use as a stair step generator

The mode: staircase generator is activated with Bluet./CCS >= 4:

  • 4: only current values increase from min: µA to max: µA with increment incr: µA
  • 5: current values rise and fall as in 4:
  • 6: Current values rise and fall as with 4: with current switching in the negative direction (staircase curves from a maximum of +4V … 0 -4V )
The dwell time per level in µs is defined via the parameters x-scale and y-scale.
 The frequency results from the length of stay and the parameter min:, incr: max: specified number of steps.

Expansion options / outlook

• larger voltage / current range Uce/Ic

For some applications the voltage range (Uce <= 12V)- or current range (<100mA) not sufficient. Operation with up to 24V (the DC/DC converter delivers a maximum of 28V) is possible without hardware changes, but at the expense of power consumption, the device should be converted to operate with 2 or more Lion cells or the NANO with 3.3 Operate V/3.6V (conversion instructions in the network, from 3.6 V safe operation of the ATMega328P with 20 MHz). Then the power consumption in standby drops drastically and you save yourself the level adjustment for the TFT display. I haven't tested the SPI interfaces to the DACs and ADCs (on the TC0372 board) at 3.6V though. A higher current range is more critical, since then, among other things, the switching regulator will also reach its limits and the TCA0372 would have to be cooled.

 • larger current constant current source / CCS

up to approx. 10 mA easily by replacing the measuring resistor (250 ohms) of the CCS to e.g. 100 ohms. Above that you probably have to change the DC/DC converter (1W/30V) and adjust the board layout and the mosfet accordingly.

  • Interface to the transit site tester (“TT”):

automatic recognition of the component including pin assignment.

This is not possible with the NANO with the now integrated functionality, since the flash memory is already 99% occupied. You would then have to use another microprocessor, e.g. the ATMega2560, which also provides a large number of port pins.
I did it - in a simpler version - with a NANO (+ port expander MCP23017 ) [7],[8] :
- Extension TT to serial interface (SoftwareSerial) and transfer Transistor type/polarity, assignment pin1,2,3 …
- Transfer (duplication) of the textool pins of the TT in a 3x3 matrix switch box,
- Installation of a serial interface in the characteristic curve recorder for reading the data from the TT.
 -.. Start the characteristic curve recorder with the control/switching data from the TT. That worked well at the time – with a considerable amount of circuitry.

• Save to SD card

Saving on an SD card is only possible with a "larger" processor (ATMega2560) see above, since the libraries require a lot of space.

• Hardware components
only the essential components / modules are listed here, the details can be found in the KiCad parts lists.

 • ILI9341 SPI display (3.2") (touch version not required)
 • Arduino NANO or clone (if possible with OptiBoot bootloader)
 • MCPM4822 dual DAC 12 bit SPI
• MCP3204 Quad ADC 12 bit SPI
 • MCP4725DAC 12bit I2C
 • ADuM1250/1 ISOLATOR I2C
 • DC/DC RH0515 1W 5V to
 • TCA0372 Dual Power Amplifier
 • LMO78-05.05 +7..+24V to – 5V
 • Textool version
 • Rotary encoder (2 pieces)
 • 10 x PhotoMos relays (e.g. TLP222A) •
 • 2 x MT3608 boards (3 V..4.2V --> 7.5V or --->15 V)
 • Lion battery incl. holder
 • Lion battery charging module
 • HM-10 modules Bluetooth 4.0

 I spent about 80 € for these components, but at the moment it is with higher ones prices to be expected.

[1] https://www.instructables.com/Transistor-Curve-Tracer/
[2] GitHub - Optiboot/optiboot: Small and Fast Bootloader for Arduino and other Atmel AVR chips
[3] Neusten Bootloader flashen auf Arduino Nano China clone – Thing King – Technik & DIY (thing-king.de)
[4] How do I install optiboot on the newest arduino IDE? - Using Arduino / Programming Questions - Arduino Forum
[5] https://chips-and-more.com/cm/pdf/K3608.pdf
[6] https://datasheetspdf.com/pdf-file/909246/AEROSEMI/MT3608/1
[7] https://www.youtube.com/watch?v=7WQflcy4CrQ
[8] https://www.youtube.com/watch?v=ybFWR0rq9jg

       Vergleich n-JFet BF256B mit p-JFet J176 Skala 6V 10mA
[9] https://www.youtube.com/shorts/95HiBvJB_0A
       NPN Skala: 2V, 5mA Parameter über Dreh-Encoder
[10] https://www.youtube.com/shorts/zspOeL5-aKM
       Kennlinienschreiber mit Präzisionskonstantstromquelle
[11] https://www.youtube.com/watch?v=6bx0Xuingm4
       Vergleich n-JFet BF256B mit p-JFet J176 Skala 1V 5mA
[12] https://www.youtube.com/shorts/LGOkdHCgLTI
        npn Test mit incr. 2,4,10,20 µA, Skala 4V 5mA
[13] https://www.youtube.com/watch?v=azDgB1QcklA&t=22s
       NPN Basistrom 5µA ... 240µA
[14] https://www.youtube.com/shorts/C6GhPL9V9Fg
      Kennlinienschreiber als bipolare Konstanstromqelle : Bereich 0 µA ... +/-1 mA
[15] https://www.youtube.com/watch?v=UbXDh1cYR88
      Kennlinienschreiber mit variablen Treppenstufengenerator
[16] https://www.youtube.com/watch?v=DciN_z7VIQE
      Kennlinienschreiber: Vergleich npn BC548 mit pnp BC558
[17] https://www.youtube.com/watch?v=N3BHx4Dii8s